corPatch Adult Therapy Electrodes 05120.5

Change Notice

The corPatch therapy electrodes have been optimised by

the manufacturer.

There may be 2 versions of the ADULT therapy electrodes circulating.

Below are the physical differences between the old and the new therapy electrodes.

New Adult

Therapy Electrodes

Rev: BA00000001

Previous Adult Therapy Electrodes

Rev: AA00000001

- Improved pictogram of ADULT therapy electrode placement, LA & LL labelling removed

- Pictogram of therapy electrode placement for paediatric patients

- Sealed pouch

- Smaller adhesive surface

- Mesh incorporated into the gel to prevent residue from remaining on the skin

- Pictogram of ADULT therapy electrode placement

- LA & LL labelling

- Zip pouch

No change in practice is necessary.

Both styles of Adult Therapy Electrodes may be used for paediatric patients over 20kg. However, we recommend using corPatch paediatric electrodes for paediatric patients up to 25kg.

*corPatch Easy Paediatric electrodes automatically limit the energy to a maximum of 100J.

See the following page for more images

corPatch Adult Therapy Electrodes 05120.5

Change Notice